Tausand

Tempico Software 1.2.0

Software for Tempico TP1004

Logotipo

Descripción generada automáticamente

User’s manual

Rev 1.1

June 30th, 2025

[www.tausand.com](http://www.tausand.com)

Contents

[1 Definitions 3](#_Toc200188532)

[1.1 Sampling time 3](#_Toc200188533)

[1.2 Edge type 3](#_Toc200188534)

[1.3 Lifetime measurements 3](#_Toc200188535)

[1.3.1 Fluorescence Lifetime measurements 4](#_Toc200188536)

[2 Description 4](#_Toc200188537)

[3 Specifications 4](#_Toc200188538)

[4 Installation 4](#_Toc200188539)

[4.1 Windows 4](#_Toc200188540)

[4.2 Ubuntu 6](#_Toc200188541)

[4.3 Mac 7](#_Toc200188542)

[5 Software operation 8](#_Toc200188543)

[5.1 System configuration 8](#_Toc200188544)

[5.1.1 Windows 8](#_Toc200188545)

[5.1.2 Ubuntu 8](#_Toc200188546)

[5.1.3 Mac 9](#_Toc200188547)

[5.2 Open the software and select a Tempico device 9](#_Toc200188548)

[5.3 Channel settings 11](#_Toc200188549)

[5.4 General settings 13](#_Toc200188550)

[5.5 Start-Stop Histogram measurements 13](#_Toc200188551)

[5.5.1 Perform a measurement 13](#_Toc200188552)

[5.5.2 Save data and plots 17](#_Toc200188553)

[5.6 Lifetime histograms 19](#_Toc200188554)

[5.6.1 Settings before a measurement 19](#_Toc200188555)

[5.6.2 Perform a measurement 20](#_Toc200188556)

[5.6.3 Fitting 21](#_Toc200188557)

[5.6.3.1 Exponential 23](#_Toc200188558)

[5.6.3.2 Shifted Exponential 23](#_Toc200188559)

[5.6.3.3 Kohlrausch 23](#_Toc200188560)

[5.6.3.4 Double Exponential 23](#_Toc200188561)

[5.6.4 Save data and plots 23](#_Toc200188562)

[5.7 Estimated Counts 25](#_Toc200188563)

[5.7.1 Settings for signal measurement 25](#_Toc200188564)

[5.7.2 Settings before measurement 26](#_Toc200188565)

[5.7.3 Perform a measurement 29](#_Toc200188566)

[5.7.4 Save data 34](#_Toc200188567)

[5.7.5 Save plot 35](#_Toc200188568)

[6 Technical documentation 36](#_Toc200188569)

[7 Drivers 36](#_Toc200188570)

[8 Technical support 36](#_Toc200188571)

# 1 Definitions

## Sampling time

The sampling time in an experiment refers to the time window during which data is collected. For example, if we count how many times the laboratory door opens over the course of 2 hours, the sampling time would be 2 hours.

## Edge type

The edge type indicates the type of transition that the time-to-digital converter (TDC) must use to start or stop a measurement. There are two main types: rise edge and falling edge. An event arriving at a TDC must be a short-duration pulse that exceeds the device's resolution. Therefore, if the transition goes from a low level (0) to a high level (1), we call it a rise edge. On the other hand, if we want to register an event when the signal transitions from a high level (1) to a low level (0), we call it a falling edge. The Tausand Tempico device allows for configuring the edge type for both the start signal and the stop signals. Below is a time diagram representation with a rise edge configuration for both the start signal (red) and the stop signal (blue).



Now we can see the representation of a *falling edge* configuration for the start signal (red) and the stop signal (blue).

## Lifetime measurements

Lifetime measurements are used to determine how long a molecule, particle, or physical system stays in an excited state before returning to its ground state, or to quantify the duration of a short-lived compound. Most of these measurements are performed by measuring the time difference between two photons.

One of the most widely used techniques for this type of measurement is time-correlated single photon counting (TCSPC). This technique relies on the emission of a pulse from a laser, which acts as a trigger to start the measurement. If the laser's frequency is correct, it partially excites the sample. Subsequently, the sample decays and emits a photon back. Some of these photons are detected by an electronic system, which converts them into electrical pulses. These electrical pulses mark the end of the measurement. The time difference between the start pulse and the stop pulse is mapped onto a histogram. Ultimately, the obtained histogram provides information about the temporal distribution of the sample's exciting state.

Tempico software is capable of capturing data in techniques that involve the emission and detection of pulses, such as in single photon counting. Several lifetime measurement methods can be applied depending on the type of analysis being performed.

## Fluorescence Lifetime measurements

This technique involves the use of a fluorophore to measure the time it takes to return to the ground state from the excited state. It is primarily used for the analysis of chemical and biological samples. The analysis is conducted by creating an environment with a fluorophore, and any change in the sample causes a modification in the fluorescent environment. In this way, temporal differences associated with interactions or changes in the sample can be obtained.

# Description

Tempico Software is designed to facilitate the use of the Tausand Tempico TP1004 devices. It is a graphical interface that allows for time histogram measurements and *lifetime* measurements. From the software, it is possible to configure the device, as well as save the measurement data and generated graphs in various text and image formats, respectively. Additionally, Tempico Software provides extra functions, such as creating custom adjustments for certain graphs, and allowing monitoring the device to ensure the measurement is being performed correctly.

# Specifications

Tempico software is designed to run on devices with low requirements, allowing any laboratory to use it. The minimum requirements for Tempico Software to function are as follows:

* **OS**: Windows 7 Home Premium, Windows 7 Professional, Windows 7 Ultimate, Ubuntu 20.04, mac OS Big Sur
* **CPU**: Intel Pentium Dual-Core E5200
* **RAM**: 1 GB
* **Storage**: 500 MB available


For measurements with large data sets or long durations (greater than 6 hours), the performance of Tempico Software may decrease. However, on higher-performance devices, this issue is mitigated.

# Installation

## Windows

The first step is to download the software from the Tausand website: <https://www.tausand.com/downloads/>. From this page, an executable file will be downloaded, which should be run to complete the installation by following the steps provided.

When opening the installer, it will request permission to run, to which we should respond affirmatively. After that, a window will appear asking us to select the installation path. We can choose the desired path; by default, the software will be installed in the Program Files folder.



Next, we will be asked if we want to create a shortcut on the desktop. To do so, we should check the corresponding box.

Interfaz de usuario gráfica, Texto, Aplicación, Correo electrónico

Descripción generada automáticamente

After completing these two steps, a tab will open showing us the installation path, allowing us to verify it before proceeding with the installation.

Interfaz de usuario gráfica, Texto, Aplicación

Descripción generada automáticamente

Once the installation process is complete, we can click "Finish". It is important to check whether we want the program to run immediately, by checking or unchecking the corresponding box.

Interfaz de usuario gráfica, Aplicación

Descripción generada automáticamente

## Ubuntu

On Ubuntu, you only need to download the compressed (.zip) file from the official page. Once downloaded, extract its contents and run the AppImage file found inside. With this process, Tempico Software will be ready to use on your system.

Interfaz de usuario gráfica, Texto, Aplicación, Word


El contenido generado por IA puede ser incorrecto.

Interfaz de usuario gráfica, Texto, Aplicación

El contenido generado por IA puede ser incorrecto.

## Mac

On macOS, you only need to download the .zip file and extract its contents to your preferred folder. The file named “Tempico Software” is the one you need to run to launch the program.



# Software operation

The software is used in the same way across all operating systems; however, the preparations for running the software may differ depending on the operating system.

## System configuration

To ensure the proper functioning of Tempico Software, it is necessary to perform certain validations on the operating system.

## Windows

You should ensure that the Tempico device is recognized by the COM ports. To do this, follow these steps: If, when connecting and turning on the device for the first time, a window appears indicating that the Tempico device is being configured, it means that Windows has recognized it correctly. If this does not happen, you can access the Device Manager without having the device connected and navigate it to the COM and LPT ports section. Then, connect and turn on the device. The port table should be updated; if it does not, close and reopen the window. When doing so, a new COM device should appear, indicating that the Tempico has been correctly detected.

## Ubuntu

**ls /dev/tty\***

In this case, it is also necessary to verify that the ports are correctly detecting the device. To do this, we will use the command above, which lists the devices connected to the serial ports. First, we will run this command without the Tempico device connected and then run it again with the device connected. We will observe which new device has been detected

**sudo adduser "username" dialout**

If an error occurs indicating that the serial port could not be opened, it is possible that the user does not belong to the dialout group. To resolve this, we will execute the command above. This will allow us to access the system's serial ports.

Si el comando no es suficiente y el software sigue lanzando errores debe verificarse cual es el puerto tty que se esta usando y debe ejecutarse lo siguiente:

**sudo ch mod 666 /dev/tty"port"**

## Mac

**ls /dev/tty.\***

To verify that the device is correctly connected to Mac, we must, with the Tempico turned off, execute the command above. This will provide a list of connected devices. Then, we run the command again with the Tempico turned on and observe which new port appears in the list.

If the device is not detected, the software will not recognize the connection, making it impossible to use the software. In this case, please contact Tausand support at: support@tausand.com.

## Open the software and select a Tempico device

When opening the software, an image with the program's logo is displayed, followed by a window that allows you to select the port to which you want to connect the device.

Logotipo

Descripción generada automáticamente

Interfaz de usuario gráfica, Aplicación

Descripción generada automáticamente

If the software is opened and the device is not turned on, the window will not recognize the connection. Therefore, it is necessary to click the "Refresh" button to update the list of ports with a Tempico device.

Once connected, the software will open with all options enabled for use. The software always opens in the histogram window.

Interfaz de usuario gráfica

Descripción generada automáticamente

If "Cancel" is selected or the dialog box is closed, the options to select a channel or start a measurement will not be available.

Interfaz de usuario gráfica

Descripción generada automáticamente

If this happens, we can click the "Connect" button again to search for devices.

Interfaz de usuario gráfica

Descripción generada automáticamente

It is also possible to disconnect the device if we want to use another Tempico. To do so, click the "Disconnect" button. If an error occurs while connecting to the device, the following error window will be displayed:

Interfaz de usuario gráfica, Texto, Aplicación

Descripción generada automáticamente

If this happens, it should be verified that no other program has established communication with the Tempico, and ensure that the USB connection, whether the cable or the physical port, is in good condition. If no issues are found after verifying this, support should be contacted to find a solution to the problem.

## Channel settings

Once the connection with the device is established, it is possible to access the settings for each channel. To do so, go to the tab at the top, select "Settings," and then choose "Channel Settings," where you can view the current settings of the device:

Interfaz de usuario gráfica, Texto, Aplicación, Correo electrónico

Descripción generada automáticamente

For each channel, you can modify the number of average cycles, the mode that defines the time range for measurements, the number of stops, the edge type and the stop mask. All these options provide a dropdown menu with the options allowed by the device (refer to the Tausand Tempico manual for more information), except for the stop mask.

Interfaz de usuario gráfica

Descripción generada automáticamente

Interfaz de usuario gráfica, Aplicación, Word

Descripción generada automáticamente

Interfaz de usuario gráfica, Aplicación

Descripción generada automáticamente

Interfaz de usuario gráfica, Aplicación, Word

Descripción generada automáticamente

For the "Stop Mask," you can enter a number between 0 and 4000 µs.



Once the desired values have been changed, click the "Apply Changes" button to transmit the changes to the device. To verify the changes, you can return to the tab and check that the values are the same as those previously set.

## General settings

The software features another window that displays the options affecting all channels of the device, including the start channel. To access it, go to the top, select "Settings," and then click on "General Settings":

Interfaz de usuario gráfica, Aplicación

Descripción generada automáticamente

From there, it is possible to set the "Threshold Voltage" and the "Number of Runs" for the measurements. The "Threshold Voltage" value allows a range between 0.90 and 1.60, with increments of 0.01, while the "Number of Runs" allows up to a value of 1000 (for more information on these settings, refer to the device manual). Once the changes have been made, click on "Save Changes," and to confirm, return to the tab and check that the values are the same as those set.

## Start-Stop Histogram measurements

"Start-stop" measurements are those that capture data from the basic functionality of the Tempico device. When a start pulse is received and, subsequently, a stop pulse arrives, the temporal difference between the two pulses is obtained. This point is plotted on a frequency histogram.

## Perform a measurement

To start the "start-stop" measurements, the channels to be sampled must be selected. To do so, check the corresponding checkboxes located on the left side of the window.

Interfaz de usuario gráfica, Aplicación

Descripción generada automáticamente

Afterward, click on "Begin Measurement" (). Depending on the number of channels selected, the corresponding graphs will be created on the right side of the window. Once the measurement begins, a histogram will be obtained for each channel.

Gráfico

Descripción generada automáticamente

It is possible to interact with the graphs generated for each channel, either by dragging the range or adjusting the zoom with the mouse wheel. The histogram is generated according to the zoom level of the plot. Therefore, by zooming in and examining the range in more detail, more precise points will be obtained, indicating where each data point falls.

Interfaz de usuario gráfica, Aplicación

Descripción generada automáticamente

Conversely, if we increase the temporal axis range, the number of bars will decrease.

Gráfico, Gráfico de líneas

Descripción generada automáticamente

Right-clicking on the graph enables additional options for interacting with it, such as manual zoom on each axis, changing the selection mode with the mouse, or exporting the image. Since these options come from the library used to generate the graph, it is not recommended to interact with the graph through these functions. However, the user is free to use them if desired.

Gráfico

Descripción generada automáticamente

The status bar indicates whether the measurement is being carried out correctly. If any channel is not receiving data, the status bar will display a message to indicate this.

Interfaz de usuario gráfica, Texto, Aplicación

Descripción generada automáticamente

If the start channel is not receiving data, the following message will be displayed:

Interfaz de usuario gráfica, Texto, Aplicación

Descripción generada automáticamente

A possible solution to this issue is configuring the threshold voltage. It is recommended to set it to 1.6 V and check if the measurement is performed correctly. If the error persists, try different values for this parameter. If the issue continues and it is determined that it is not a fault in the setup or the input signal, Tausand support should be contacted.

Start-stop histograms are designed to take measurements indefinitely. Once the experiment is considered complete, click the "End Measurement" button ( ) to stop data collection. If you wish to clear the data during the experiment, there's no need to stop the measurement. At the bottom of the left menu, there are "Clear" buttons to erase both the data and current graphs. Four buttons are available to clear data for each channel independently ( , , , ).

## Save data and plots

If the user wishes to save the graphed data, they should click the save button ( ), which will open a menu allowing them to select the format in which to save the files.

Interfaz de usuario gráfica, Aplicación, Tabla

Descripción generada automáticamente

After selecting the desired format, click the accept button ( ), which will display a dialog box indicating that the file has been saved successfully

Interfaz de usuario gráfica, Texto, Aplicación

Descripción generada automáticamente

Files are saved in all operating systems in the "TempicoSoftwareData" folder within "Documents." By default, this folder is automatically created when the program is opened for the first time.

The data is saved in the same format. First, it includes the configuration applied when the data was collected, followed by the raw histogram information. This allows the user, if desired, to later manipulate the data according to their needs.

Texto

Descripción generada automáticamente

If the user wishes to save the graphs generated by the software during the measurement, they must click on the "Save Plots" button ( ). A window will then open, allowing them to select the image format.

Interfaz de usuario gráfica, Texto, Aplicación

Descripción generada automáticamente

After selecting the format, click the "Accept" button ( ). As with the previous case, a dialog box will appear indicating that the image has been successfully saved. The image will be saved in the same location as the text files, across all operating systems.

Interfaz de usuario gráfica, Texto, Aplicación

Descripción generada automáticamente

The saved image is the one displayed in the software at the moment of clicking "Accept." Therefore, the zoom applied to the graph will also be reflected when saving the image. If multiple channels are selected for measurement, the 4 images will be saved separately.

Texto

Descripción generada automáticamente

## Lifetime histograms

To record lifetime measurements based on the time difference between input pulses and emitted excitation pulses, Tempico Software provides multiple tools to fulfill this functionality.

## Settings before a measurement

Unlike the measurement of start and stop histograms, the preliminary setup for a lifetime measurement is a bit more elaborate. First, we must select the start channel and the stop channel for our measurements.

If a channel other than start is selected to initiate the measurement, a periodic electrical signal must be added to the start channel, as Tempico cannot measure signals on stop channels without a registered start event. This configuration risks losing data that arrives before the periodic signal is emitted. The pulse period must be short enough to reduce data loss. Tempico does not register stops occurring within less than 12 ns, making this setup useful for mitigating such effects.

After selecting the channel, the bin width must be chosen, representing the histogram's bar width. For example, with a bin width of 12 ns, if there are two measurements at 1 ns and 5 ns, both will contribute to the same histogram bar's frequency. The number of bins prevents histogram saturation and usually limits the time axis range. For instance, selecting a bin width of 480 ps and 50 bins results in a 24 ns range. Unlike start-stop histograms, lifetime measurements are designed with a limit, allowing the number of measurements to be specified. A measurement is counted when both start and stop channels record a value.

Interfaz de usuario gráfica, Texto, Aplicación

Descripción generada automáticamente

It's important to clarify that the configurations for the number of runs, number of stops, stop mask, and mode have no effect on this measurement. By default, the program sets the number of runs to 100, the number of stops to 1, the stop mask to 0, and the mode according to the maximum time parameter. If the maximum time exceeds 500 ns, the measurement operates in mode 2; otherwise, it operates in mode 1.

## Perform a measurement

Once the desired configuration is selected, click the start button ( ) to begin the measurement. Like start-stop histograms, the program displays a status bar indicating the measurement's progress. If any channel is not recording, it will be shown. In this case, the percentage of the measurement is also displayed, as the number of desired measurements can be specified.

Imagen de la pantalla de un computador

Descripción generada automáticamente con confianza baja

Interfaz de usuario gráfica, Texto, Aplicación

Descripción generada automáticamente

The program also provides additional information about the measurement status, including the total number of starts, the total number of complete measurements (i.e., start and stop), and the total elapsed time in HH:MM:SS format.

Interfaz de usuario gráfica, Texto

Descripción generada automáticamente

## Fitting

Con el fin de obtener más información acerca de una medición, el usuario puede realizar un fit sobre la gráfica. Al realizar el fit obtiene en la tabla de la esquina superior derecha información más detallada acerca de los parametros de la ecuación a ajustar. Obtiene sus valores, su desviación estándar y sus respectivas unidades.

Gráfico

Descripción generada automáticamente

Gráfico

Descripción generada automáticamente

In some cases, reliable values may not be estimated, and the adjustment will indicate "nan" for those parameters. The quality of the fitness depends on the initial parameters set before the estimation. Therefore, users are free to modify these parameters. They can also revert to the default values provided by the software.

Interfaz de usuario gráfica, Texto, Aplicación, Correo electrónico

Descripción generada automáticamente

The software offers various types of adjustments:

## Exponential

For fluorescence decay processes and certain radioactive processes, it is common to encounter exponential distribution. This distribution involves two parameters: by default, the initial value of ​ is set to the maximum found on the graph, and ​ corresponds to the average of the measured times.

## Shifted Exponential

In some cases, an exponential distribution may exhibit a shift due to noise or inherent system characteristics. For these situations, an option for an exponential distribution with a shift is provided, allowing the user to correct the shift and obtain a pure exponential distribution. The initial parameters ​ and ​ are the same as those used by default in the exponential distribution. However, the additional parameters and are set to 0 by default.

## Kohlrausch

In the case of fluorescence lifetime measurements, interactions may occur both in the environment and with the fluorescent material, which can be interpreted as noise in the final measurement. To obtain a profile that represents the dynamics of fluorescence, a Kohlrausch fit can be used. The default initial value for the parameter is 1.

## Double Exponential

In some cases, the data may have more than one component, and the user may wish to separate the parameters of each. For these cases, a double exponential fit can be used. The default values are the same as those used in the exponential distribution; the default value for is 1, and the value for is the same as .

## Save data and plots

Like the histogram functionality, the user has the option to save both the data and the image. If the user clicks the **Save Data File** button ( ), a window will appear to select the desired format for saving the data. The user can choose between txt, csv, or dat formats.

Interfaz de usuario gráfica, Aplicación

Descripción generada automáticamente

If the user clicks the **Accept** button ( ), a window will appear confirming that the file has been successfully saved and displaying the path where it was stored.

Interfaz de usuario gráfica, Texto, Aplicación

Descripción generada automáticamente

If the user clicks **Save Plot** ( ), they can choose from three formats: PNG, TIFF, and JPG. Upon clicking the **Accept** button, a window will confirm that the file has been successfully saved and display the path where it was stored.

Interfaz de usuario gráfica, Texto, Aplicación

Descripción generada automáticamente

If an adjustment is applied, both the image and the text data will include the equation of the adjustment made, along with the calculated parameters.

Tabla

Descripción generada automáticamente

Gráfico, Histograma

Descripción generada automáticamente

## Estimated Counts

Tempico Software does not directly provide the exact count of pulses received on each channel, since the Tempico device accepts a maximum of five incoming pulses after a start pulse. However, the software offers a feature that enables reliable count estimation by analyzing the time intervals between pulses.

## Settings for signal measurement

To estimate pulse counts in Tempico Software, it is essential that the measurement includes a periodic start pulse of very short duration, and that the source to be measured is connected to one of the available channels (A, B, C, or D). In this process, software settings such as average cycles, mode, number of stops, and stop mask have no effect, as the system automatically configures these parameters during the measurement. The software calculates the average time differences between the pulses detected on the active channel, then computes the inverse of this value to determine the pulse frequency per unit of time, and finally converts the result into seconds, obtaining an estimate of the number of pulses per second. A graph is shown below to illustrate a representative signal used for this type of estimation.

Diagrama

El contenido generado por IA puede ser incorrecto.

The signal clearly shows the periodic start pulse that must be emitted, along with two sets of pulse measurements recorded on Channel A (blue) and Channel B (red), respectively. For Channel A, the software identifies the times t1 Channel A and t2 Channel A. To obtain t1, it subtracts the timestamp of pulse 1 from pulse 2, and for t2, it subtracts pulse 2 from pulse 3. Then, the software calculates the average between these two values, obtains their inverse, and converts the result into seconds. For Channel B, the same procedure is followed: the average is calculated between t1 Channel B, t2 Channel B, t3 Channel B, and t4 Channel B, then the inverse is obtained, and the result is converted to seconds. This final value is reported as the measurement result.

The Tempico device has a 4-millisecond window to receive a stop pulse after a start pulse is received. Additionally, it is essential that at least two pulses are detected after the start pulse during the measurement process, as otherwise it will not be possible to calculate the time differences between pulses. Therefore, the input signal must have a minimum frequency of 500 pulses per second for the software to accurately estimate the counts.

## Settings before measurement

For this functionality, users can configure the channels on which the count estimation will be performed, as well as choose from different types of graphical displays. The merge option overlays the curves of all channels on a single graph; separated creates an individual graph for each channel; and detached opens each graph in a separate, resizable dialog window. Additionally, the detached table option displays the temporal count history in an independent dialog. The time range parameter allows users to define the time intervals over which the measurement graph will be displayed.

It is important to note that not all settings become locked once the measurement begins. The only configuration that becomes restricted is channel selection, which prevents enabling channels that were not selected prior to starting the measurement. All other settings can be adjusted dynamically while the measurement is running. The following figure shows the configuration panel for these options.

Interfaz de usuario gráfica, Texto, Aplicación, Correo electrónico

El contenido generado por IA puede ser incorrecto.

Once the user starts the measurement, the system automatically determines the number of stops required to begin count estimation. This step is critical, as configuring the Tempico device with several stops greater than the actual number of pulses received after a start pulse would result in no valid measurements being detected. To prevent this issue, the software analyzes the incoming signal behavior and automatically adjusts the appropriate number of stops before starting the counting process.

Imagen que contiene Interfaz de usuario gráfica

El contenido generado por IA puede ser incorrecto.

If the software detects that a channel is unable to register more than 2 stops, it will notify the user and prompt whether to stop the measurement or continue with the remaining active channels. This mechanism ensures user control over the process and helps prevent unreliable results from channels with insufficient signal.

Interfaz de usuario gráfica, Texto, Aplicación

El contenido generado por IA puede ser incorrecto.

If the user chooses to continue the measurement, the channel for which no valid data was found will be automatically disabled and cannot be reselected unless the current measurement is stopped and a new one is started. This ensures that only channels with valid signals are considered during the estimation process.

Interfaz de usuario gráfica, Texto, Aplicación, Correo electrónico

El contenido generado por IA puede ser incorrecto.

If the user chooses not to continue, the software will automatically terminate the measurement, ensuring that no channels without valid signals are processed.

## Perform a measurement

To start the measurement, you must click the Start Measurement ( ) button, which will begin the stop number estimation process described earlier. Once this process is complete, the measurement will begin to run correctly.

First, we have the graphical section; if the Merge Graphics ( ) option is selected, all curves from the selected channels will be displayed in a single chart.

Interfaz de usuario gráfica, Gráfico

El contenido generado por IA puede ser incorrecto.

Whereas if the Separate Graphics ( ) option is selected, the curves for each channel will be displayed individually, each in its own separate chart.

Gráfico, Gráfico de líneas

El contenido generado por IA puede ser incorrecto.

We also have the Detached Graphics ( ) option, which displays the graphs in separate dialog windows that open independently from the main window.

Gráfico, Gráfico de líneas

El contenido generado por IA puede ser incorrecto.

Time filters can be applied to the graphs using the Time Range dropdown menu. This allows you to adjust the time range on the X-axis to view the measurement history in detail. For example, if you select 50 seconds, all graphs will display the last 50 seconds of the recorded data.S

Imagen que contiene Tabla

El contenido generado por IA puede ser incorrecto.

Gráfico, Tabla, Excel

El contenido generado por IA puede ser incorrecto.

During the measurement, there is also a table section that displays the history of the estimated count values along with the corresponding hour, minSSutes, and seconds of registration. This section is located below the graphs displayed on the screen.

Interfaz de usuario gráfica, Aplicación

El contenido generado por IA puede ser incorrecto.

If the user wants to view the table more clearly, they can enable the *detached table* option available in the settings panel. This will display the table in a separate dialog window, showing the same data as the one at the bottom of the main window, but with enhanced detail.

Interfaz de usuario gráfica

El contenido generado por IA puede ser incorrecto.

Tabla

El contenido generado por IA puede ser incorrecto.

Likewise, in the bottom-left corner of the main window, the user can view the value and the uncertainty corresponding to the most recently reported measurement.

Interfaz de usuario gráfica, Texto, Aplicación

El contenido generado por IA puede ser incorrecto.

As before, the user can choose to view the current values in detached mode by clicking the corresponding checkbox.Interfaz de usuario gráfica

El contenido generado por IA puede ser incorrecto.

After that, a separate dialog will open displaying the current measurement values.

Interfaz de usuario gráfica, Aplicación, Word

El contenido generado por IA puede ser incorrecto.

If a channel stops emitting values during the measurement process, the software will notify the user through the status bar. Additionally, the results table will display the message Low Counts, indicating that the device is unable to perform the measurement because the number of received counts is insufficient.

Interfaz de usuario gráfica, Aplicación, Correo electrónico

El contenido generado por IA puede ser incorrecto.

Interfaz de usuario gráfica, Aplicación

El contenido generado por IA puede ser incorrecto.

Once measurements are registered again on the channel, the process will resume normally.

## Save data

The user is responsible for deciding when to stop the measurement. To do this, they must click the Stop Measurement ( ) button. After this, the Save Data ( ) button will be enabled, allowing the user to save the data in plain text format. Upon clicking this button, a window like those used in other functionalities will appear, prompting the user to choose the format in which they wish to save the measurement.

Interfaz de usuario gráfica, Aplicación, Correo electrónico

El contenido generado por IA puede ser incorrecto.

Among the available formats for saving the data are .txt, .csv, and .dat. Upon clicking Accept, a window will appear confirming that the values have been successfully saved and displaying the path where the files are located.SS

Interfaz de usuario gráfica, Texto, Aplicación

El contenido generado por IA puede ser incorrecto.

The file will store both the timestamp of the recorded data and the corresponding estimated count value.

Imagen que contiene Gráfico de dispersión

El contenido generado por IA puede ser incorrecto.

## Save plot

Similarly, once the measurement is completed, the option to save the plots () will be enabled. As in the previous sections, a window will appear allowing the user to select the format in which they wish to save the graph, including PNG, TIFF, and JPG.

Interfaz de usuario gráfica, Texto, Aplicación

El contenido generado por IA puede ser incorrecto.

After clicking the accept button, a dialog window will open indicating the path where the image was saved, including the file name and selected format.

Interfaz de usuario gráfica, Texto, Aplicación

El contenido generado por IA puede ser incorrecto.

The graph will be saved exactly as it appears on screen. If the time range filter has been applied, the image will reflect the corresponding values on the X-axis. Additionally, if the graph has been visually adjusted—either stretched or displayed with specific width-to-height proportions—those proportions will be preserved when the graph is saved.

Gráfico, Gráfico de líneas

El contenido generado por IA puede ser incorrecto.

## Technical documentation

Tempico Software is open source, allowing access to the software's source code and developer documentation through the GitHub platform. It also provides a guide to the code and the creation of respective installers for each operating system.

<https://github.com/Tausand-dev/TempicoSoftware>

## Drivers

Your computer may automatically download and install from the Internet the required drivers once you connect your Tausand Tempico device to a USB port. If you require to install them manually, they are available at <https://www.tausand.com/downloads>.

## Technical support

For technical support, contact us at [support@tausand.com](mailto:support@tausand.com).

Visit our website to check for the latest documentation and software releases.

[www.tausand.com](http://www.tausand.com)